## Cust Rqmts
Team #: | P00000 | Team Name: | Sample MSD Project
Date: | 2018-01-31 14:50 | Document Owner: | Team Member
Revision #: | 1
Customer Rqmt. # | Importance | Description | Comments/Status
CR1
CR2
Cust. Rqmt. #: enables cross-referencing (traceability) with engineering requirements
Importance: 1=must have, 2=nice to have, 3=preference only.
Description: organize as primary and secondary requirements (hierarchy) -- Ulrich exhibit 4.8
Comment/Status: allows tracking of questions, proposed changes, etc; indicate if you are meeting the requirement ("met") or not ("not met")

## Eng Rqmts
Team #: | P12016 | Team Name: | Navigation Aid
Date: | 2018-01-31 14:50 | Document Owner: | Ernst Killander
Revision #: | 1
rqmt. # | Importance | Source | Function | Engr. Requirement (metric) | Unit of Measure | Marginal Value | Ideal Value | Actual Value | Test Result | Direction of Improvement | Comments/Status | Test Number
S1 |  | CR 2.1 | Identify current orientation | Error in compass heading | degrees | 10 | 1 |  |  | down
S2 |  | CR 2.1 | Identify current map location | Error in RSSI from RFID | % | 10 | 5 |  |  | down
S3 |  | CR 2.2 | Identify current map location | Distance between RFID tags | m | 12 | 8 |  |  | down
Rqmt. #: enables cross-referencing (traceability) and allows mapping to lower level requirements within separate documents
Source: Customer requirement #, regulatory standard (eg. EN 60601), and/or "implied" (must exist but doesn't have an associated customer rqmt), constraint
Description: quantitative, measureable, testable details
TEST RESULT KEY |  |  |  |  | DIRECTION OF IMPROVEMENT KEY
X | Does not meet expectation |  |  |  | X | Must be within a specified range.
Δ | Caution-Undetermined if specification is met |  |  |  | ↑ | Improvement occurs with increasing test value.
O | Meets specification |  |  |  | ↓ | Improvement occurs with decreasing test value.

## Cust vs Eng Rqmts
Team #: | P00000 | Team Name: | Sample MSD Project
Date: | 2018-01-31 14:50 | Document Owner: | Team Member
Revision #: | 1
CR1
CR2
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0

## Sample Test Plan
Team #: | P00000 | Team Name: | Sample MSD Project
Date: | 2018-01-31 14:50 | Document Owner: | Team Member
Revision #: | 1
Subsystem/ Function/ Feature Name: |  |  | Volume Sensitivity Testing
Date of Test: | 2013-11-05
Performed By: __Steve DiGerardo, Paulina Klimkiewicz, Matt Burkell and David Herdzik___
Tested By: __Steve DiGerardo, Paulina Klimkiewicz, Matt Burkell and David Herdzik___
Concluded Condition of meeting Engineering Specification |  |  |  | Ο- Meets Marginal Value
I. TESTING SPECIFICATION
Specification Number | Importance | Source | Function | Specification (Metric) | Unit of Measure | Marginal Value | Ideal Value | Comments/Status
S1 | 1 | PRP | System | Volume Control | Liters | 0.2 ± 0.2
II. EQUIPMENT REQURIED
Specification Number | Equipment or Instrumentation required
S1 | Fully Charged Battery, Pump, Differential and Pressure Sensor Unit, Digital Mass Flow sensor
III. DATA COLLECTION STRATEGY
Specification Number | Data acquisition strategy
S1 | The goal of this test is to characterize our sensors to the appropriate flow rate as seen by the digital mass flow rate sensor. By calibrating our sensor unit to the sensitivity of the digital flow sensor, our venturi system design will be able to provide that sensitivity when in operation.
III. TESTING FLOWCHART
IV. RAW DATA ACQUISITION
Volume Testing: System Configuartion with stop watch testing Using a .7 time multiplier | Flow Rate (l/s) | Flow rate (l/s) | Time (s) | Volume (l) | Volume Comanded (l) | % Diff
| 15 | 0.25 | 3.9 | 0.975 | 1.02 | -4.412%
| 15 | 0.25 | 4 | 1 | 1.02 | -1.961%
| 15 | 0.25 | 3.9 | 0.975 | 1.02 | -4.412%
| 15 | 0.25 | 4.1 | 1.025 | 1.02 | 0.4902%
| 30 | 0.5 | 4 | 2 | 2.03 | -1.478%
| 30 | 0.5 | 4.1 | 2.05 | 2.03 | 0.9852%
| 30 | 0.5 | 4.1 | 2.05 | 2.03 | 0.9852%
| 30 | 0.5 | 4 | 2 | 2.03 | -1.478%
V. RESULTS
VI. CONCLUSION
In this test, we are confirming the accuracy of our breath delivery system to ensure it meets volume sensitivity. In this test, at different flow rates, we are comparing the amount of flow given by the pump to the amount of flow as read by a digital sensor. From the test points above we test a lower range flow of 15 L/min and a higher range flow of 30 L/min. In each of the flow domains we are able to read similar data to that of the digital mass flow sensor. Our maximum error experienced is 4%. This is well under the 10% sensitivity flow desired. Thus we can conclude that our pump performance and sensing capabilites are able to sense the the flow accurately.
Gather all necessary materials.
Reduce flow rate to lower limits
Flow volume through the Mass Flow Rate Sensor.
Record and gather all data for the control volume sensitivity over full operating spectrum of the flow rate.
Increase flow rate with pump in increasing intervals
Remove Mass flow rate sensor and hook up Differential and static sensor unit to the pump.
Reduce flow rate to lower limits
Flow volume through the differential and static sensor unit .
Record and gather all data for the control volume sensitivity over full operating spectrum of the flow rate.
Increase flow rate with pump in increasing intervals (same intervals as control)
Plot both both volume sensitivities as flow rate in the pump is increased.
Coorelate and calibrate sensor unit to match control sensitivity.
S1 Test Plan

## S1 Test Plan
Team #: | P00000 | Team Name: | Sample MSD Project
Date: | 2018-01-31 14:50 | Document Owner: | Team Member
Revision #: | 1
Subsystem/ Function/ Feature Name:
Date of Test:
Performed By:
Tested By:
Concluded Condition of meeting Engineering Specification:
I. TESTING SPECIFICATION
Specification Number | Importance | Source | Function | Specification (Metric) | Unit of Measure | Marginal Value | Ideal Value | Comments/Status
S1
II. EQUIPMENT REQURIED
Specification Number | Equipment or Instrumentation required
S1
III. DATA COLLECTION STRATEGY
Specification Number | Data acquisition strategy
S1
III. TESTING FLOWCHART
IV. RAW DATA ACQUISITION
V. RESULTS
VI. CONCLUSION
S1 Test Plan